## Slide 1
National and International Timing
  Per Olof Hedekvist
  RISE Research Institutes of Sweden

## Slide 2: Commonly used Acronyms
BIPM – Bureau International des Poids et Mesures
GUM – Guide to the Expression of Uncertainty in Measurement
UTC –Coordinated Universal Time
  UTC(k) – local realization of UTC
TAI – International Atomic Time
MJD – Modified Julian Date
  # of days since Sputnik launch
  1/1 2020 = MJD 58849
GNSS – Global Navigation Satellite System
  GNSS CV – Common View.

## Slide 3: Time & frequency signal / data - Requirements
Better silent than wrong
Traceable to common reference
  Known uncertainty
Multiple sources must be weighted with respect to stability
  Historical data determines weight.

## Slide 4: Time & frequency signal / data - Traceability
reference
user
Unbroken chain of information
Bidirectional
  Link-data not sufficient, client clock data is necessary
  Synch to GNSS signal not traceable

## Slide 5: UTC
International time since 1972
680 clocks in 88 labs (2018)
  Local realizations UTC(k)
Continuous comparison, data collected by BIPM
  Calculates weighted average
  Sends correction data to all participants
  Local timescales use data for improved UTC(k) realizations

## Slide 6: RISE (former SP)
Department of Measurement Science and Technologies
National Metrology Institute in Sweden
PNT-group
    5 members
    3 FET
UTC(SP) since 1/4 1996
Main site in Borås, additional site in Stockholm since 2013
T&F transfer using GNSS CV
TWSTFT station in Borås
Redundant fiber connection through SUNET
  10 GBE channel in 100 GBE wavelength.

## Slide 7: Cs and H-masers
25 clocks reported to BIPM
  8 H-masers (+2 in Sept. 2019)
  15 Cs high performance
  2 Cs standard performance

## Slide 8
Secure clock site
Construction of clock labs in cavern
A dual Time and Frequency lab has been built in secured cavern
  Stockholm area
  > 15 m below surface level
  EM-shielded, vibration absorbers, shock wave protection
  UPS, batteries and 3 diesel generators
  Climate control, with redundancy
  Locations divided in two separate, identical labs
  Separate redundant power and cooling supply.
  3 Cs and 1 active HM in each lab
    5071A-001 and MHM2020
  Identical set of timing and communication equipment in both

## Slide 9: RISE National Timescale UTC(SP)
BIPM
UTC(SP) – GPS/TWSTFT
Clock 1 SP
Clock 2 SP
Clock N SP
UTC(SP)
Correction
UTC(SP) – UTC
f(KL) – f(TAI)
GPS/TWSTFT
Group-clock
Timeinterval-
counter
UTC(SP) - Clock

## Slide 10: Historical deviation from UTC
Last 5 years within ± 10 ns from UTC
Standard deviation = 3,48 ns

## Slide 11: Traceable to UTC(SP)
Stockholm nodes generates timescales traceable to UTC(SP)
  BRG1 and BRG2
Continuous comparison, daily data log.
  BRG1: 2,63 ns offset, 0,77 ns STDEV (1σ)
  BRG2: 1,72 ns offset, 0,33 ns STDEV (1σ)
  Based on 80 days of data summer 2019

## Slide 12: Communication and dissemination
Both sites connected to SUNET
3 ntp-servers in Borås
2 ntp-servers in Stockholm
Dissemination to external customers
  Telecom operators
  Defense
  Financial market
  Governmental authorities
WR under deployment for wider use

## Slide 13: Communication and dissemination
Netnod
  5 Internet exchange nodes + one testnode
  2xCs at each site generates timescale traceable to UTC(SP)
  Internal communication can generate timescale based on 12 Cs if disconnected from external communication
  Dissemination to wider range of external customers
    NTP
    PTP
    NTS in development

## Slide 14: National Research Projects
Time transfer in SDH-networks
  Passive frame detection
  Operational in 10 Gb/s
  Bandwidth efficient
  Customized hardware
  Tested in SUNET
    Extended to MIKES
  2005-2015
Polarization effects in fiber transmission
  Evaluation of effects in polarization dependent receivers
  50 Hz + harmonics

## Slide 15: Temperature variations
Fiber influenced by temperature even when deployed in ground
Ground temperature (1 m depth) varies up to ± 10 °C over the year
  Data from USDA, Acriculture Services
  Largest variations in subarctic regions
Fiber deployed in air affected by daily temperature variations as well as fast movement by wind and rain.

## Slide 16: Environmental influences
Data from time transfer in Sweden – Finland
Borås – Stockholm, red curve
  Deployed in ground
  8 router stations above ground
Stockholm – Espoo, green curve
  Deployed at bottom of Baltic Sea
  No routers along line
  Discrepancies caused by human interference at endnodes.

## Slide 17: Polarization
Coherent transmission requires mixing with an optical LO at receiver.
  Very polarization sensitive process
  Datacom splits polarizations, detects both, and extract data through DSP.
  Polarization variation at n*50 Hz, when fiber is deployed near power lines
  Standard DSP introduces unknown variable delay

## Slide 18: Limitations in duplex fibers
Unknown asymmetries limits time transfer performance
Constant asymmetries must be identified through alternative calibration technique
Asymmetric variations must be estimated through additional parameters

## Slide 19: Limitations in simplex fibers
Maximum distance limited by attenuation
  Cannot be fully compensated by amplification
  Special installation in networks, not generally applicable
Any duplex sections must be within climate control
  Path lengths as identical as possible
  Asymmetry in wavelengths

## Slide 20: EU Metrology research
Article 185 in EU research program
Fixed amount available to each country each year.
NMI and external partners in collaboration

## Slide 21: EU Fiber Research Projects
EURAMET EMRP/EMPIR
  NEAT-FT (2012 – 2015)
    TF in fiber between UTC(SP) and UTC(MIKE)
  OFTEN (2016 – 2019)
    Evaluation of real time TF in SUNET max distance and applied for TF to OSO
  TiFOON (2019 – …)
  WRITE (2018 – …)
Aim to connect NMIs and users in Europe with fiber for time and frequency dissemination

## Slide 22: Fact sheet
Kick-off | 1 June 2019
Duration | 3 years
Budget | Approx. 2 M€
Consortium NMI, Industry, NREN, Academic | 15 partners Coordinated by NPL
JRP 18SIB06
TiFOON - TIme and Frequency Over Optical Networks
Advanced time/frequency comparison and dissemination through optical telecommunication networks

## Slide 23: Overview
Aim
To advance fibre-based frequency transfer capabilities in Europe towards a universal tool for time and frequency metrology and beyond.
Motivation
Redefinition of the SI second Fibre links enable the remote clock comparisons, but reliability and availability are insufficient.
International time scales Fibre-based time transfer offers performance and resilience beyond satellite techniques.
Economical sustainability through combined optical frequency and time services in shared telecom fibre.
Applications beyond metrology Fibre links will provide access to ultra-stable and accurate optical time and frequency references.
Vision ≠ Reality

## Slide 24: Wider Impact
Enhanced international time scales
Satellite navigation (e.g. Galileo)
Telecommunication networks
All-optical time keeping
Optical frequency standards, oscillators and signalling
Femtosecond timing?
Pan-European fibre network for T&F
Dissemination of state-of-the-art frequency standards
Secure traceable timing
Global challenges
Climate – Water – Energy

## Slide 25: EU Galileo Research Projects
GSOp
  operation of GST (Galileo Time Scale).
  One of five UTC-lab supplying UTC(SP) GNSS and TW-data
GRC-MS
  monitoring quality and accuracy of GST through local receivers
3PfD
  Studies on Galileo PRS-signals

## Slide 26: Time’s up
per.olof.hedekvist@ri.se